## Requirements Register
ID | Date Added | Requirement Description | Extra Details | Source | Mentioned in Report? | Satisfied? | Critical to Safety | Functional Rating | Stakeholder Rating | Final Rating | Priority | Status
REQ-001 | 2023-08-01 | The design shall produce a DC output voltage of 0 - 200V | Increments of 0.05V | Project Brief | Y | Yes | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-002 | 2023-08-01 | The design shall produce a DC output voltage that is hardware interlock limited to 100V |  | Project Brief | Y | Yes | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-003 | 2023-08-01 | The design shall produce an AC output voltage of 0 - 200Vrms |  | Project Brief | Y | Yes | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-004 | 2023-08-01 | The design shall produce an AC output voltage with hardware interlock limiting it to 50 Vrms. | Hardware can still mean software-controlled | Project Brief | Y | Yes | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-005 | 2023-08-01 | The design shall produce a maximum output current of 10mA |  | Project Brief | Y | Yes | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-006 | 2023-08-01 | The design shall have an adjustable output maximum current cutoff/trip limit tuneable in the range of 1-10mA. | Adjustment of the trip limit, but it is not current-limiting | Project Brief | Y | Yes | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-007 | 2023-08-01 | The design shall protect the output from overcurrent by switching off the output in under 10 milliseconds. |  | Information Session | Y | In progress | Yes | 5 | 5 | 20 | Safety Critical | Approved
REQ-008 | 2023-08-01 | The design shall include a suitable, sealed case for safe handling. |  | Project Brief | Y | In progress | Yes | 2 | 4 | 16 | Safety Critical | Approved
REQ-009 | 2023-08-01 | The design shall include a suitable case of a self-extinguishing material. | UL-94 | Project Brief |  | In progress | Yes | 2 | 4 | 16 | Safety Critical | Approved
REQ-010 | 2023-08-01 | The design shall undergo preliminary testing with output voltages not exceeding 30V (peak) for both DC and AC. | TBC - Not client or design requirement, UWA enforced? | Project Brief | Y | No | Yes | 1 | 5 | 16 | Safety Critical | Approved
REQ-011 | 2023-08-06 | The design shall comply with all relevant legal and regulatory requirements. |  |  |  | In progress | Yes | 2 | 4 | 16 | Safety Critical | Approved
REQ-012 | 2023-08-06 | The design shall consist of a visual indication light to show when the interlock is engaged or disengaged. |  |  | Y | In progress | Yes | 1 | 4 | 15 | Safety Critical | Approved
REQ-013 | 2023-08-01 | The design shall utilise a computer or microcontroller. |  | Project Brief | Y | Yes | No | 5 | 5 | 10 | Critical | Approved
REQ-014 | 2023-08-01 | The design shall utilise a physical interface to take in user inputs. | Keyboard or touchscreen or knobs or buttons | Project Brief | Y | In progress | No | 5 | 5 | 10 | Critical | Approved
REQ-015 | 2023-08-01 | The design shall produce an AC output frequency of 50-300kHz sine wave. |  | Project Brief | Y | Yes | No | 5 | 5 | 10 | Critical | Approved
REQ-016 | 2023-08-01 | The design shall produce a switchable output 100% amplitude modulated signal with a square wave signal for AC & DC outputs. | Reword | Project Brief | Y | Yes | No | 5 | 5 | 10 | Critical | Approved
REQ-017 | 2023-08-01 | The design output connection shall be a BNC or high-voltage BNC. |  | Information Session |  | Yes | No | 5 | 5 | 10 | Critical | Approved
REQ-018 | 2023-08-10 | The design shall have one output channel. | Need to confirm is one or more outputs (Will not need AC and DC voltage outputted at the same time) | Information Session |  | Yes | No | 5 | 5 | 10 | Critical | Approved
REQ-019 | 2023-08-01 | The design shall utilise a combination of a display and LEDs to indicate the status of the device. |  | Project Brief | Y | In progress | No | 4 | 5 | 9 | High | Approved
REQ-020 | 2023-08-01 | The design shall allow for output voltage adjustments with steps of 0.05V or less over full range of voltages. |  | Project Brief | Y | Yes | No | 4 | 5 | 9 | High | Approved
REQ-021 | 2023-08-01 | The design shall allow for output frequency adjustments with steps of 10kHz for the carrier frequency. |  | Project Brief | Y | Yes | No | 4 | 5 | 9 | High | Approved
REQ-022 | 2023-08-10 | The design shall allow for current limit adjustments with steps of 2mA. |  | Information Session | Y | Yes | No | 4 | 4 | 8 | High | Approved
REQ-023 | 2023-08-01 | The design shall produce a square wave signal with frequency range of 10-100Hz. |  | Project Brief |  | Yes | No | 4 | 4 | 8 | High | Approved
REQ-024 | 2023-08-01 | The design shall source input power from an external power supply adapter. |  | Project Brief | Y | Yes | No | 3 | 5 | 8 | High | Approved
REQ-025 | 2023-08-01 | The output type (DC/AC), voltage, modulation (ON/OFF) and output enable shall all be controlled via a digital input. |  | Information Session | Y | Yes | No | 3 | 5 | 8 | High | Approved
REQ-026 | 2023-08-10 | The design shall minimise the voltage ripple as much as possible. | Especially at low frequencies (below 100kHz) as it can damage MEMS. | Information Session |  | Yes | No | 3 | 5 | 8 | High | Approved
REQ-027 | 2023-08-01 | The design shall allow for output frequency adjustment with steps of 10Hz for the square wave. |  | Project Brief |  | Yes | No | 3 | 4 | 7 | High | Approved
REQ-028 | 2023-08-01 | The design shall accept nominal input voltages in the range of 5 - 30VDC. | Prefer 12VDC | Project Brief | Y | Yes | No | 2 | 5 | 7 | High | Approved
REQ-029 | 2023-08-01 | The design shall include a suitable case that allows for programming and computer-based control without disassembly. | Remove computer-based control? | Project Brief |  | In progress | No | 2 | 4 | 6 | Medium | Approved
REQ-030 | 2023-08-01 | All design files shall be uploaded to a Git repository with an open source / open hardware license for the project partner and other researchers. | Preferred CERN-OHL-P standard for hardware and Apache 2.0 for software | Project Brief |  | No | No | 1 | 5 | 6 | Medium | Approved
REQ-031 | 2023-08-01 | The prototyping activities shall not exceed $350 for reimbursement purposes. |  | Project Brief | Y | Yes | No | 1 | 5 | 6 | Medium | Approved
REQ-032 | 2023-08-06 | The design shall utilise components that are available in sufficient supply . |  |  | Y | Yes | No | 2 | 4 | 6 | Medium | Approved
REQ-033 | 2023-08-10 | The design shall operate nominally in room temperature conditions. |  | Information Session |  | Yes | No | 3 | 3 | 6 | Medium | Approved
REQ-034 | 2023-08-10 | The design shall have an expected lifetime of approximately 10 years. |  | Information Session |  | In progress | No | 3 | 3 | 6 | Medium | Approved
REQ-035 | 2023-08-01 | The design shall be implemented on a PCB or protoboard if time constrained. |  | Information Session | Y | Yes | No | 2 | 3 | 5 | Medium | Approved
REQ-036 | 2023-08-01 | The design shall minimise electromagnetic interference produced as far as is practicable. | CE compliant? Nic to investigate bars around box. | Information Session | Y | In progress | No | 2 | 3 | 5 | Medium | Approved
REQ-037 | 2023-08-01 | The design's size and weight shall be as low as possible. | Testable through weight tallying and comparing to alternative components. | Information Session |  | Yes | No | 1 | 2 | 3 | Low | Approved
REQ-038 | 2023-08-08 | Frequency distortion of the sine wave shall be ≤ 2% | i.e. Total Harmonic Distortion (THD). Requires testing. | Information Session | Y | In progress | No | 3 | 3 | 6 | Medium | Approved
REQ-039 | 2023-08-08 | Step voltage ripple should be at least 10x smaller than the minimum step size (5 mV) | Requires testing. | Information Session |  | In progress | No | 3 | 3 | 6 | Medium | Approved

## Rating Matrix
Safety Rating (1-5)
1: Missing this requirement results have no impact on the health and safety of any individuals
2: Missing this requirement has a very low risk of minor injury
3: Missing this requirement has a moderate risk of minor injury OR a very low risk of major injury
4: Missing this requirement has a high risk of minor injury OR a moderate risk of major injury
5: Missing this requirement has a high risk of major injury
Cost Rating (1-5)
1: Missing this requirement results in no financial cost
2: Missing this requirement results in minor financial cost
3: Missing this requirement results in moderate financial cost
4: Missing this requirement results in major financial cost
5: Missing this requirement results in major financial cost for multiple stakeholders
Stakeholder Rating (1-5)
1: This requirement does not benefit any stakeholders
2: This requirement has slight benefit to some stakeholders
3: This requirement has major benefit to some stakeholders OR slight benefit to important stakeholders
4: This requirement has major benefit to important stakeholders
5: This requirement has been specifically listed by important stakeholders
Functional Rating (1-5)
1: This requirement does not affect the functionality of the design
2: This requirement may affect the functionality of the design
3: This requirement will affect the functionality of the design
4: This requirement is critical to the functionality of the design
5: This requirement is critical to the functionality of the design & successful design delivery
Critical to Safety
Yes: Not meeting this requirement will create a hazard to person/s
No: Not meeting this requirement will not create any hazard

## Categories
Final Rating Range | Priority Categories | Priority Number | Description | Status Categories | Description | Verification Method Categories | Description | Member Name Categories | Description | Source Categories | Description
0 | Low | 1 | Any requirement that is not key to the functional requirements of the design and is considered low importance to client satisfaction. | Proposed | This status indicates that the requirement has been identified and suggested but has not yet been approved or started. | Inspection | Non-destructive examination of a product using one or more of the 5 senses.--Review the design schematics to determine if they meet the requirements. | LB | Liam Bettles | Project Brief
4 | Medium | 2 | Any requirement that is not key to achieving the functional requirements of the design and is considered high importance to client satisfaction. | Approved | This status indicates that the requirement has been reviewed and accepted by the stakeholders and can be included in the project plan. | Demonstration | To conduct a trial run (simulation) of the project and observe if it can achieve the basic functionality as designed. | AC | Anders Christensen | Information Session
7 | High | 3 | Any requirement that is key to achieving the required functionalities and must be included for client satisfaction | Rejected | This status indicates that the requirement has been evaluated and deemed unsuitable or unnecessary and will not be pursued. | Test | Simulate general working conditions and check whether they meet the specific requirements of stakeholders. | NC | Nicholas Curwood-Wagner | Technical Query
10 | Critical | 4 | Any requirement that is key to achieving the functional requirements and must be included for a successful design delivery to occur | In Progress | This status indicates that work has started on the requirement and is currently being developed or implemented. | Analysis | Simulate specific working conditions, continuously monitor the relevant indicators of the project to ensure strict compliance with the standards. | PP | Pete Philip Paul
Non-compliance can create a hazard (regardless of rating) | Safety Critical | 5 | Any requirement that creates a hazard if it is not met. | Verified | This status indicates that the requirement has been fully developed and has undergone testing or quality assurance, and meets the criteria established for its success. |  |  | TW | Thomas Wilkinson
|  |  |  | Cancelled | This status indicates that the requirement has been terminated and will no longer be pursued. |  |  | JZ | Jason Zagari
|  |  |  | On Hold | This status indicates that the requirement is temporarily paused and will not be pursued until further notice or a certain condition has been met. |  |  | GRP | Whole Group